MOLECULAR KARYOTYPING: TOWARDS IMPROVED PRE- AND POSTNATAL GENETIC DIAGNOSIS

Joris Robert Vermeesch1, Cindy Melotte1, Guy Froyen2, Steven Van Vooren4, Binita Dutta3, Nicole Maas1, Stefan Vermeulen5, Björn Menten5, Frank Speleman5, Bart De Moor4, Paul Van Hummelen3, Peter Marynen2, Jean-Pierre Fryns1 and Koen Devriendt1.

1Center for Human Genetics, University Hospital Gasthuisberg, Leuven, Belgium

2Center for Human Genetics, Flanders Interuniversity Institute for Biotechnology (VIB4), Department of Human Genetics, Leuven, Belgium

3MicroArray Facility, Flanders Interuniversity Institute for Biotechnology (VIB), Leuven, Belgium.

4 ESAT-SISTA K.U. Leuven, Belgium

5Department of Medical Genetics, Ghent University, Ghent, Belgium

# Abstract

Array CGH enables the identification of chromosomal copy number changes.  The availability of clone sets covering the human genome opens the possibility for the widespread use of array CGH for both research and diagnostic purposes.  In this manuscript we report on the parameters that were critical for successful implementation of the technology, explore the essential quality criteria and discuss the potential benefits and pitfalls of the technology for improved pre- and postnatal diagnosis. We propose to term the genome wide array CGH “molecular karyotyping”, in analogy with conventional karyotyping that use staining methods to visualise chromosomes.

# Introduction

One of the ultimate aims of pre- and postnatal cytogenetic testing has been to provide as complete as possible a genome-wide analysis for the detection of chromosome aneuploidies and segmental aneusomies. Traditionally, karyotyping has performed this role. In addition to its clinical importance, the identification of a chromosomal aberration in specific patients has proven to be a successful way to identify the implicated genes and gain insight in the pathogenesis of different genetic conditions.

Over the last decades, the progressively higher resolution of the cytogenetic techniques has lead to an increasingly larger proportion of recognizable chromosomal aberrations in patients with congenital anomalies and/or mental retardation (MR). However, the resolution of traditional cytogenetic techniques appears limited to ~5 Mb and smaller chromosomal aberrations often remain hidden. The identification of submicroscopic subtelomeric alterations in 3-7% of idiopathic mental retardation patients (Flint et al., 1995; Knight et al., 1999; Slavotinek et al 1999; Riegel et al., 2001; Rosenberg et al., 2001) as well as the sporadic reports of submicroscopic interstitial chromosomal rearrangements suggests that a substantial portion of ideopathic mental retardation may be caused by smaller chromosomal rearrangements.  These observations make it clear that higher resolution screening techniques for the detection of small deletions or duplications at any chromosomal position will drastically increase the elucidation of human genetic diseases.

Fluorescence in situ hybridization (FISH) has significantly increased the resolution and made possible the detection of smaller chromosomal aberrations. Molecular probes derived from specific chromosomal regions are fluorescently labeled and hybridized on metaphase or interphase chromosome spreads from the patients cells. Deletions or amplifications of the probe region can subsequently be detected by the loss or gain, respectively of signals that are expected for a normal diploid individual. This technology has very strong diagnostic value, but is too laborious as a genomic screening tool for unknown rearrangements.

More recently, so called array or matrix CGH utilizes mapped DNA sequences in a microarray format as a  platform for the detection of chromosomal deletions/duplications. In this technique, genomic DNA from the patient is labeled with one fluorescent dye while a normal reference sample is labeled with a different dye and these samples are co-hybridized to the array containing the genomic DNA targets. Chromosomal imbalances across the genome can thus be quantified and positionally defined by analysing the ratio of the fluorescence of the two dyes along the targets. The resolution of array CGH depends on the size of the genomic fragments as well as their density.  Proof of principle was established in 1997 (Pinkel et al.1997; Solinas-Toldo et al., 1997).  Since then, only few laboratories mastered the technology, mainly for the detection of chromosomal amplifications in cancers (Solinas-Toldo et al., 1997; Pollack et al., 1999; Albertson, 2000; Bruder et al., 2001).  Chromosomal duplications or deletions were initially technically challenging to detect (Carter et al., 2002).  Very recently, several other groups have successfully increased the sensitivity to the point that single copy changes can reliably be detected, which has led to the first reports of genome wide screens in both human (Pollack et al., 1999; Snijders et al., 2001, Vissers et al., 2003; Shaw-Smith et al., 2004; Schaeffer et al., 2004; Schoumans et al., 2004) and mouse (Hodgson et al., 2001).   However, due to these technical challenges, the technique has not made the transition to clinical practice.   In this manuscript we explore some of the technical aspects essential for optimal results in array CGH and we inquire the quality criteria an array experiment should reach to be reliable in a clinical setting.

# Materials and methods

## Array CGH

Arrays were constructed using a 1 Mb Clone Set, which contains 3527 BAC and PAC clones (Fiegler et al., 2003).  DNA was isolated from 1 ml bacterial culture using Millipore genomic DNA purification kit following the instructions of the manufacturer (Millipore, Billerica, Mass). BAC and PAC DNA obtained after purification was reconstituted in storage buffer supplied by the manufacturer. The BAC/PAC DNA was amplified by “degenerate oligo nucleotide primed PCR” (DOP PCR) using a modification of the protocol described by Fiegler et al. (2003). Pre-amplification was performed in a 50µl reaction mixture containing PCR Buffer (20 mM Tris HCl (pH 8.4), 50 mM KCl), 5mM MgCl2 , 200 µM dNTP’s (Amersham Pharmacia Biotech, Piscataway, NJ), 2 µM primermix (5’-CCGACTCGAGNNNNNNCTAGAA-3’; 5’-CCGACTCGAGNNNNNNTAGGAG-3’;5’-CCGACTCGAGNNNNNNTTCTAG-3’), 2.5 units Platinum Taq DNA polymerase (Invitrogen, Carlsbad, CA) and 2-10 ng template DNA.  Reactions were performed using the GeneAmp PCR System 9700 thermocycler (Perkin Elmer, Applied Biosystems,----) using the following conditions: 3 min denaturation  at 94°C followed by 9 thermal cycles with denaturation for 1.5 min at 94°C, annealing for 2.5 min at 30°C, subsequently ramping with temperature increment of 0.1°C/s to 72°C and a 3 min extension at 72°C.  Additionally 30 cycles were performed consisting of 1 min at 94°C, 1.5 min at 62°C and 2 min at 72°C, followed by a final elongation step at 72°C for 8 minutes.

In a second round of amplification an aminolinked primer was used.  A 100 µl reaction was performed by combining 2 µl of the first reaction product with 1.5 µM primer (5’-NH2–GGAAACAGCCCGACTCGAG-3’), 200 µM dNTP’s (Amersham Biosciences), PCR buffer, 5 mM MgCl2  and 2.5 units  Platinum Taq DNA polymerase (Invitrogen). Thermal cycles were performed as follows: 10 min denaturation at 95°C; 35 cycles of 1 min at 95°C, 1.5 min at 60°C and 7 min at 72°C and a final elongation at 72°C for 10 minutes.  Following amplification, the PCR products were purified by Qiaquick 96 well PCR purification kit following the manufacturers instruction (Qiagen Inc., Valencia, USA). The purified DOP PCR products were EtOH precipitated.

DOP PCR products were spotted on either Code Link Bioarray System slides (Amersham Biosciences, Piscataway, NJ), type VII star silane coated mirror slides (Amersham Biosciences) or UltraGAPS amino-silane-coated slides (Corning, NY).  For spotting on amino silane coated slides, regular DOP PCR primers (without aminolinker) were used to amplify the BAC DNA (Van Buggenhout et al., 2004).  Purified DOP PCR product was reconstituted in 20 µl of 80% DMSO solution containing nitrocellulose (0.37 µg/ml) at an average concentration of 250 ng/l. The products were arrayed on using a Molecular Dynamics Generation III printer (Amersham Biosciences). Before hybridization, the slides were humidified, dried and target DNA was cross-linked in a UV-Stratalinker with 50 mJ UV light (Stratagene, Amsterdam, The Netherlands). For spotting on Code Link Bioarray System slides, 5’ aminolinked DOP PCR products were dissolved in 250 mM sodium phosphate buffer (pH 8.5) containing 0.001 % sarkosyl.  Products were spotted at a concentration of 200ng/µl with a Molecular Dynamics Generation III printer (Amersham Biosciences). The clones were printed in duplicate. After printing, slides were pretreated by incubating them in a humidified chamber saturated with NaCl solution for 24 hours. Next day, slides were treated with 1%  NH4OH  solution on shaker followed by 0.1% SDS for 5 minutes. The slides were rinsed in water, placed in 950C waterbath for 2 minutes  and then at – 200C for 1 min. The slides were rinsed again in H2O and finally dried by centrifugation. These slides are kept under dehydrated conditions until further use.

To test the printing quality, the pretreated slide was hybridized with 71.1 pmoles of Cy3 labeled degenerate  oligo primer (5’-Cy 3 –GGAAACAGCCCGACTCGAG-3’) for  1 hour. Following hybridization, the slide was washed with H2O and scanned to check quality of printing.

DNA from an anonymous cell line with karyotype 46,XX or 46,XY was used as a reference. Test and reference gDNA was labeled by a random prime labeling system (Bioprime DNA Labeling System, Invitrogen) using Cy3 and Cy5 labeled dCTP’s (Amersham Biosciences) as described (Van Buggenhout et al., 2004). The labelling efficiency was checked with the Nanodrop ND-1000 spectrophotometer (Nanodrop Technologies, Rockland, DE). DNA concentration was determined by measuring the intensities at 260 and 280 nm.  Cy5 and Cy3 incorporation was measured at respectively 650 and 550 nm.  For each probe the corresponding specific activity was calculated by the following formula:  (total ng of probe x 103 )/ (molecular weight of nucleotide times the total pmol of dye incorporated). Except for some small modifications, probe preparation and pre-blocking of the slide was performed as described by Fiegler et al.(2003).  In short, equal amounts (200pmol) of Cy5 and Cy3 labeled probe were combined with 200 µg Cot-1 DNA, followed by an ethanol precipitation.  Resuspension of the pellet was done in hybridization buffer (50% formamide, 10% dextran sulphate, 0.1% Tween 20, 2X SSC, 10mM Tris HCl pH 7.5) containing 400 µg yeast tRNA to hybridize a spotting area of 24x60mm.

The slide was blocked with 50 µg Cot-1 DNA and 300 µg salmon sperm DNA, dissolved in 60 µl hybridization buffer. Blocking solution and probe mixture were denaturated for 10 minutes in a water bath at 75°C.  Blocking solution was then placed on the slide covered with a coverslip (24 x 60 mm) and placed in a humid chamber.  Meanwhile the probe was placed at 37°C for pre-annealing.  After 1 hour of blocking, the coverslip was removed and probe was placed on the slide.  After placing a coverslip (24 x 60 mm), the slide was placed in a humid chamber saturated with 20% formamide and 2X SSC.  Hybridization was allowed to take place for 2 nights at 37°C.   While optimizing the protocol, it was noticed that the targets at the outer ends of the slide often show reduced signal intensities when sealing the slide with cement rubber.  Eliminating the sealing of the cover-slips by  hybridizing the slides in small humid chambers yielded equal intensity ratios over the entire (full) slide.  The hybridization was allowed to take place for 2 nights under a coverslip in a humid chamber saturated with 20% formamide and 2XSSC.  Post hybridization washes were performed as described by Fiegler et al. (2003).  In short , the cover-slip was removed by placing the slide in PBS with 0.05% Tween 20,  followed by a 10 minutes wash in a fresh solution of  PBS /Tween 20  at room temperature, 30 minutes in 50% formamide/ 2 x SSC at 42°C and 10 minutes in 1 x PBS/0.05% Tween 20 at room temperature.  Slides were spin dried at 1000 rpm for 5 minutes.

Three types of slide were compared: amino-silane-coated slides on which the target DNA is cross-linked by UV light, Code Link Bioarray System slides on which the DNA is covalently linked via the amino group of the DOP primer and type VII star silane coated mirror slides. When hybridized with 250 pmol probe each, the signal-to-noise ratios on amino-silane-coated and Code Link Bioarray System slides were respectively 6-10 and 20-30 for the Cy3 and Cy5, respectively. The mirror slides were hybridized with 200 pmol probe resulting in signal-to-noise ratios of respectively 5.3 and 2.7 for the Cy3 and Cy5. These values are derived from at least 10 hybridizations.

##### Image and data analysis

Arrays were scanned at 532nm (Cy3) and 635nm (Cy5) using the GenIII scanner (Amersham) for the mirror slides or the Agilent G2565BA MicroArrayScanner System (Agilent Inc., Palo Alto, CA) for the CodeLink and UltraGAP slides. Image analysis was done using ArrayVision (Imaging Research Inc, St Catharines, Ontario, Canada).  All further data analysis was performed with Excel (Microsoft Inc.). Spot intensities were corrected for local background and only spots with signal intensities at least twofold above background signal intensities were included in the analysis.  For each clone a ratio of Cy5 to Cy3 fluorescent intensity was calculated.  Normalization of the data was achieved by dividing the fluorescent intensity ratio of each spot by the mean of the ratios of the autosomes.  Finally the normalized ratio values of the duplicates were averaged and a log2 value was calculated. Datapoints for which the variation among the two intensity ratios was larger than 10% were excluded from the analysis.

# Results

# Parameters influencing array CGH quality

The two main technical challenges when performing array CGH are obtaining  adequate signal to noise ratios and low standard deviations of the intensity ratios.  In contrast to expression arrays, genomic array CGH deals with a vastly more complex probe mixture and the copy number variations are much lower. In pre- and postnatal diagnostic samples copy number ratios vary mainly between ½  for deletions or 3/2 for duplications.   To obtain the best possible signal to noise ratios, controlling the fluorescent dye incorporation during labelling, assaying probe quantity, the type of slides used and adequate hybridisation skills proved essential.  Table 1 provides an overview of these different parameters and lists the measures that were taken to optimize these parameters.

Since usually only a small subset of the genome is expected to be aberrant in pre- and postnatal constitutional abnormalities, the intensity ratios of patient versus control DNA will be equal for most spots on the array.  This offers a unique possibility to measure the quality of an array experiment by determining the standard deviation of all intensity ratios.  The log2 values of these intensity ratios behave in a linear fashion since the relation of the intensities of both dyes inversely influence the value of the intensity ratio at a spot.  The normalised log2 of the intensity ratio values will typically fluctuate around 0.   The lower this fluctuation, the better the array CGH experiment.  To obtain the lowest standard deviation of the intensity ratios, adequate signal-to-noise ratios are essential as well as an optimised Cot-1/probe ratio (figure 1). During these optimisation experiments we noted that not only insufficient Cot-1 (figure 1c) but also an excess of Cot-1  (figure 1a) leads to an increase in the SD.   The standard deviation can be further reduced by averaging several independent hybridisation experiments (figure 2).

### Quality criteria and threshold values

When karyotyping, the resolution of an experiment is defined by the number of bands that can be discerned.  When performing a genome wide array CGH experiment, in theory, all clones will provide a value and the resolution of the experiment is defined by the number of genomic fragments on the array as well as by their size.  However, in practise, a number of spots fail due to bad quality, e.g. bad spotting or hybridisation efficiency.  Therefore, it is essential to report the percentage of spots that were amenable to analysis. At present in our hands, typically over 97.5% of the spots provided adequate intensity ratio values, i.e. signal-to noise over 2 and the SD among the duplicates less than 10%.

When performing array CGH the main challenge is to define a threshold value above which no false positives are retained without eliminating true positives (i.e. avoiding false negatives). Since the log2 transformed normalized intensity ratios fluctuate in a Gaussian fashion around 0, the standard deviation can be used to define thresholds.  A threshold level is described as mean plus or minus three or four times the standard deviation (Schwaenen et al., 2004; Shaw-Smith et al., 2004).  Using three times the standard deviation as a cut-off level 99.7% of the fragments will fall within the normal range.  This would result in about 10 false positive clones on an array with 3500 different loci.  Using four times the standard deviation as a cut-off level 99.9936 % of the fragments fall within the normal range, resulting in 1 false positive every 4 analyses. Therefore, for an array containing 3500 loci,  four times the standard deviation is the threshold value of choice.

In addition, using 4 times the SD as a threshold value defines an important array quality value: the value of the standard deviation of the log2 transformed normalized ratios.  To be workable, the value of four times the standard deviation needs to be below the detection limit of an autosomal duplication, which can be defined as the log2 transformed mean intensity ratio of the duplicated loci minus two times the standard deviation or 4*SD≤log2(3/2)-2*SD or SD≤0.096

## Polymorphisms

While 4 times the SD seems to be an adequate cut-off level based on statistical grounds, this reasoning assumes a perfect Gaussian distribution of intensity ratios and neglects the biological and experimental causes of non-Gaussian behaviour of a (small) subset of genomic fragments.  It is likely that some biological causes of non-Gaussian behaviour can lead to false positives.

Polymorphisms can be identified on arrays as those clones for which the intensity ratios of independent experiments repeatedly fall outside the above defined cut-off level. Identifying and reporting these polymorphic clones is thus an essential first step toward proper interpretation of array CGH data. When abnormal values are obtained in a number of separate experiments either using DNA from individuals without an obvious clinical phenotypic or in different experiments with DNA from individuals with very different clinical phenotypes, it can be assumed that these clones are polymorphic.  Based on the DNA analysis of 30 individuals, aberrant signal intensity ratios were obtained in at least three different experiments in 18 loci (table 2). However, their status as polymorphic loci awaits further experimental ascertainment.

Polymorphic loci will not only be detected by intensity ratios repeatedly surpassing the 4 times SD threshold value, but, when the polymorphisms are smaller than a single clone on the array, they will be evidenced as smaller deviations from the average intensity ratio. Indirect evidence for such polymorphisms comes from experiments with a low  SD.   A low SD results in a low threshold level which in turn results, contradictory, in many more DNA fragments that are (false) positive for a deletion or duplication than in a less optimal experiment with a higher SD (data not shown).

Therefore, to eliminate the further analysis of these “false positive” values, we suggest to use not only the log2(1) ± 4*SD as a threshold but, in addition, use the log2 of the mean of the duplicated loci minus two times the SD as a second higher threshold.  97.72% of duplicated loci and ~100% of deleted loci will surpass this higher threshold.  We propose that in order to trigger confirmatory experiments, the intensity ratio at a single locus has to surpass this higher threshold.  However, in case for two flanking clones the intensity ratios surpass the lower threshold, further experiments are warranted to confirm or disprove the suspected chromosomal anomaly.

To empirically test these theoretical figures, DNA from a normal cell line was hybridized versus DNA from a cell line trisomic for chromosome 13 (figure 3).  The observed log2 transformed mean of the intensity ratios of duplicated loci was 0.53 rather than the theoretical 0.58.  The SD of the log2 transformed intensity ratios at all normal spots was 0.08 while the SD of the duplicated loci was 0.1, rather than the theoretical SD of 0.08. From a total of 102  chromosome 13 derived loci, 88 were above the higher threshold, 10 between the higher and lower threshold and  4 loci below the lower threshold.  Hence,  the intensity ratios of 13% rather than the expected 2.3% of the duplicated loci are below the higher threshold value.  That more than expected values are below the threshold values is in part caused by the lower empiric mean value of the duplicated loci.  This lower mean value could  be caused by incomplete blocking or repetitive sequences and by the occurrence of some clones that contain low copy repeats which will render the theoretical value for a deletion lower than 3/2.

## Mosaicisms

## When using DNA from patients carrying mosaic segmental or chromosomal aneuploidies using array CGH, none of the datapoints derived from the aneuploid loci may reach the above defined threshold levels.   However, a series of intensity ratios will deviate from the mean value in a similar direction.  Visual inspection of array CGH data can rapidly recognize the presence of non random spreading of a series of flanking datapoints. DNA from a 100% trisomy 13 cell line DNA and a mixture containing only 20% of the trisomy 13 cell line DNA was hybridised versus DNA from a normal cell line.  In figure 3b, the presence of a trisomy 13 in 20% can be identified by viewing the intensity ratio values of all the chromosome 13 loci which are all unidirectionally above the mean intensity ratio value.   Molecular karyotyping can thus, in analogy with conventional karyotyping, readily be used to identify mosaicisms.

## Discussion

Since both genome wide array CGH analysis and conventional karyotyping, based on staining chromosomes, aim to identify chromosomal aberrations by screening the genome, we propose to call this technology molecular karyotyping.  Molecular karyotyping is likely to become part of the routine genetic diagnosis and to replace in part current karyotyping technologies.  Advantages over conventional karyotyping include a higher resolution, direct mapping of aberrations to the genome sequence, amenability to automatisation and quality control procedures and, probably, higher throughput and shorter reporting times.  Essential for the successful introduction of the technology in current cytogenetic laboratories are (1) robust protocols and a good understanding of the critical technical factors (2) quality criteria which define a successful array experiment and (3) reporting guidelines to enable the correct interpretation of the results obtained by different laboratories which includes a good knowledge of polymorphic loci in the human population.

1. Parameters influencing array CGH quality

While array formats may differ in that they may use BACs (Solinas- Toldo et al., 1997; Pinkel et al. 1997), cDNAs (Pollack et al., 1999), oligo’s (Lucito et al., 2003) or PCR fragments (Mantiprada et al., 2004) as targets, the main steps involved in an array experiment are similar and include array production,  probe preparation, hybridisation and washing and data analysis.  When comparing different protocols that have been described, a large variation in the details of the method can be discerned (Carter et al., 2002).  While all aspects of the protocol are important for optimal array CGH data, we identified the DNA/Cot-1 ratio as well as the efficiency of fluorescent dye incorporation as the main causes of inter-experimental variation. Batch-to-batch variability of commercial Cot-1 preparations has been pinpointed before as the cause of variability among different experiments (Carter et al., 2002).  The fluorophore incorporation efficiency (specific activity) has a direct effect on the sensitivity of an experiment and is thus an obvious key element for a successful array experiment.

2. Quality criteria:

In analogy with conventional karyotype, a molecular karyotype should have a well defined quality.   First, the number of clones on the array with successful hybridisations should be reported.  Second, the standard deviation of the log2 transformed intensity ratios should not exceed 0.096.  Third, a minimum threshold can be defined as plus or minus four times the standard deviation of all the log2  transformed intensity ratios.   Finally, a higher threshold value was defined  as log2(3/2)-2*SD.   If the intensity ratio of only a single loci surpasses the lower threshold it should also surpass the higher threshold to trigger confirmatory experiments.  However, if two flanking loci surpass the lower threshold, this equally triggers confirmatory experiments.

3.Current pitfalls for pre- and postnatal diagnosis: polymorphic loci

Variations in chromosome lengths and banding are well known to the cytogenetic community.  Common chromosomal variants have been documented over the last 30 years and practitioners have a good knowledge of the most common benign variants. Novel benign variants are still being uncovered.   Polymorphisms may extend several megabases (and thus multiple array loci) (Hand et al., 2000; Martin et al., 2002; Starke et al., 2003).  Not surprisingly, at the higher resolution level obtained by molecular karyotyping, similarly polymorphic loci are detected   and, due to the higher resolution, the number of variants that are observed has equally increased.  The first reports using whole genome array CGH in constitutional cytogenetic analyses have identified respectively 2/20 and 5/50 single clone anomalies in patients  which were also present in their phenotypically normal parents (Vissers et al., 2003; Shaw-Smith et al., 2004). During the analysis of 30 patients we estimated a total of 18 loci (or 0.5% of all loci on the array) within our clone set  to be polymorphic as defined by intensity ratios in different independent experiments which are concordant with a deletion or duplication of a locus. However, proof that these loci are truly polymorphic sites awaits confirmation.  In addition, it is likely that smaller intraclonal variations and polymorphisms  also exist. Albertson and Pinkel (2003) reported on the variation of the intensity ratios of a single spot derived from DNA  on chromosome 6 containing the apolipoprotein.  They suggest this variation is due to the known polymorphisms within this gene. Indirect evidence that more intraclonal polymorphisms exists comes from our experiments with very low  SD.   In such optimal experiments the low cut-off levels result in many more DNA fragments that are (false) positive for a deletion or duplication than in a less optimal experiment with a higher SD.  However, such non-Gaussian behaviour of certain spots could also have other causes such as suboptimal printing or hybridization efficiencies.

Polymorphic clone information is likely to become integrated in the genome annotations.  However, in the absence of a large scale concerted effort, the question can be raised how such polymorphic sites in the genome will be ascertained.  It is likely that continued feed-back from a series of dedicated laboratories may lead to a validated database of candidate or proven genomic polymorphisms.

#### Future prospects

While we defined molecular karyotyping as a genome wide array CGH experiment no resolution has been defined. The minimum resolution of a molecular karytype should equal but preferentially surpass the resolution obtained by conventional karyotyping,  The maximum resolution that can be obtained using BAC or PAC based clone arrays would be a genomic tiling path array of approximately 32K, which has already been achieved (Ishkanian et al., 2004).  Even higher resolutions can be obtained by using clones with overlapping BAC/PAC clones (Ishkanian et al., 2004), smaller insert fragments (Bruder et al., 2001), PCR products (Mantipragada et al., 2004) or oligonucleotides (Lucito et al., 2003).  However, molecular karyotyping at higher resolution will likely await clinical use.  First, based on the statistical principle that 4x SD will define the cut-off at  99.9936% , still 2 spots would result in false positives.  One possibility is to increase the cut-off to 5xSD or even 6xSD, which will require more stringent quality criteria for an array CGH experiment.  In addition, considering an estimated degree of 0.5-1% polymorphic clones in the genome, the identification and validation of these polymorphic loci will be a daunting task. Possibly, novel improved approaches to identify polymorphic loci may resolve this issue.  Finally, low copy repeats, often containing small copy number variations, make up 5-10% of the genome.  Aberrant ratios at these loci will be more difficult to identify.

Molecular karyotyping is likely to replace in part current karyotyping technologies based on staining chromosomes both in pre- and postnatal diagnosis.  However, besides the many advantages, molecular karyotyping has also some caveats as compared to the conventional karyotyping e.g. it fails to identify translocations and ploidy variations.  Therefore, it seems likely that we will keep on enjoying the view of banded chromosomes in the foreseeable future and that banded chromosomes will remain an invaluable tool in the genetic diagnostic laboratory.

## Acknowledgements

We would like to thank the Mapping Core and Map Finishing groups of the Wellcome Trust Sanger Institute for initial clone supply and verification.  This work was made possible by grants G.0200.03 from the FWO and OT/O2/40 from the University of Leuven.

Table 1: Parameters influencing signal to noise and standard deviation of the intensity ratios

|  | Action taken | Remarks |
| --- | --- | --- |
| ## Signal to noise |  |  |
| Labeling efficiency | Measure specific activity | Values in the range of 1 fluorophore every 30-80 bp |
| Probe concentration | Test different probe concentrations | Currently 200 or 250 pmol probe |
| Slide type | Test different slides |  |
| DNA concentration in targets | Measure DNA concentration following PCR |  |
| ## Equal hybridisation efficiency across the slide | - No sealing of the coverslips - Open hybridisation chambers | (Fiegler et al., 2003) |
| ## Standard deviation |  |  |
| Cot-1/probe ratio | Determine optimal ratio | Currently 200 pmol probe versus 100 µg Cot-1 |

Table 2: Clone list with suspected polymorphic loci*

| Clone ID | chromosome |
| --- | --- |
| RP5-1108M17 | 1 |
| RP11-294I20 | 2 |
| RP6-22D12 | 3 |
| RP11-24H13 | 4 |
| RP11-6C14 | 4 |
| RP11-17M8 | 8 |
| RP11-356M23 | 8 |
| RP11-292F22 | 10 |
| RP11-114K7 | 11 |
| RP11-161M6 | 16 |
| RP11-274A17 | 16 |
| RP11-342F21 | 17 |
| RP11-374N3 | 17 |
| RP11-537N4 | 19 |
| CTA-299D3 | 22 |
| RP6-14C6 | ND |
| RP11-99C10 | ND |
| RP1-225D2 | ND |

*excluding clones on the X and Y chromosome

ND = not determined

References

1. 	Martin, C.L., D.J. Waggoner, A. Wong, S. Uhrig, J.A. Roseberry, J.F. Hedrick, S.D. Pack, K. Russell, E. Zackai, W.B. Dobyns, and D.H. Ledbetter. 2002. "Molecular rulers" for calibrating phenotypic effects of telomere imbalance.  J Med Genet JID - 2985087R 39:734-740.

2. 	Ishkanian, A.S., C.A. Malloff, S.K. Watson, R.J. DeLeeuw, B. Chi, B.P. Coe, A. Snijders, D.G. Albertson, D. Pinkel, M.A. Marra, V. Ling, C. MacAulay, and W.L. Lam. 2004. A tiling resolution DNA microarray with complete coverage of the human genome.  Nat Genet JID - 9216904 36:299-303.

3. 	Vissers, L.E., B.B. de Vries, K. Osoegawa, I.M. Janssen, T. Feuth, C.O. Choy, H. Straatman, d. van, V, E.H. Huys, A. van Rijk, D. Smeets, C.M. van Ravenswaaij-Arts, N.V. Knoers, d.B. van, I, P.J. de Jong, H.G. Brunner, A.G. van Kessel, E.F. Schoenmakers, and J.A. Veltman. 2003. Array-based comparative genomic hybridization for the genomewide detection of submicroscopic chromosomal abnormalities.  Am J Hum Genet JID - 0370475 73:1261-1270.

4. 	Snijders, A.M., N. Nowak, R. Segraves, S. Blackwood, N. Brown, J. Conroy, G. Hamilton, A.K. Hindle, B. Huey, K. Kimura, S. Law, K. Myambo, J. Palmer, B. Ylstra, J.P. Yue, J.W. Gray, A.N. Jain, D. Pinkel, and D.G. Albertson. 2001. Assembly of microarrays for genome-wide measurement of DNA copy number.  Nat Genet 29:263-264.

5. 	Carter, N.P., H. Fiegler, and J. Piper. 2002. Comparative analysis of comparative genomic hybridization microarray technologies: report of a workshop sponsored by the Wellcome Trust.  Cytometry JID - 8102328 49:43-48.

6. 	Schaeffer, A.J., J. Chung, K. Heretis, A. Wong, D.H. Ledbetter, and M.C. Lese. 2004. Comparative genomic hybridization-array analysis enhances the detection of aneuploidies and submicroscopic imbalances in spontaneous miscarriages.  Am J Hum Genet JID - 0370475 74:1168-1174.

7. 	Lucito, R., J. West, A. Reiner, J. Alexander, D. Esposito, B. Mishra, S. Powers, L. Norton, and M. Wigler. 2000. Detecting gene copy number fluctuations in tumor cells by microarray analysis of genomic representations.  Genome Res JID - 9518021 10:1726-1736.

8. 	Mantripragada, K.K., I. Tapia-Paez, E. Blennow, P. Nilsson, A. Wedell, and J.P. Dumanski. 2004. DNA copy-number analysis of the 22q11 deletion-syndrome region using array-CGH with genomic and PCR-based targets.  Int J Mol Med JID - 9810955 13:273-279.

9. 	Fiegler, H., P. Carr, E.J. Douglas, D.C. Burford, S. Hunt, C.E. Scott, J. Smith, D. Vetrie, P. Gorman, I.P. Tomlinson, and N.P. Carter. 2003. DNA microarrays for comparative genomic hybridization based on DOP-PCR amplification of BAC and PAC clones.  Genes Chromosomes Cancer JID - 9007329 36:361-374.

10. 	Pollack, J.R., C.M. Perou, A.A. Alizadeh, M.B. Eisen, A. Pergamenschikov, C.F. Williams, S.S. Jeffrey, D. Botstein, and P.O. Brown. 1999. Genome-wide analysis of DNA copy-number changes using cDNA microarrays.  Nat Genet JID - 9216904 23:41-46.

11. 	Hodgson, G., J.H. Hager, S. Volik, S. Hariono, M. Wernick, D. Moore, D.G. Albertson, D. Pinkel, C. Collins, D. Hanahan, and J.W. Gray. 2001. Genome scanning with array CGH delineates regional alterations in mouse islet carcinomas.  Nat Genet 29:459-464.

12. 	Albertson, D.G. and D. Pinkel. 2003. Genomic microarrays in human genetic disease and cancer.  Hum Mol Genet JID - 9208958 12 Spec No 2:R145-R152

13. 	Pinkel, D., R. Segraves, D. Sudar, S. Clark, I. Poole, D. Kowbel, C. Collins, W.L. Kuo, C. Chen, Y. Zhai, S.H. Dairkee, B.M. Ljung, J.W. Gray, and D.G. Albertson. 1998. High resolution analysis of DNA copy number variation using comparative genomic hybridization to microarrays.  Nat Genet 20:207-211.

14. 	Bruder, C.E., C. Hirvela, I. Tapia-Paez, I. Fransson, R. Segraves, G. Hamilton, X.X. Zhang, D.G. Evans, A.J. Wallace, M.E. Baser, J. Zucman-Rossi, M. Hergersberg, E. Boltshauser, L. Papi, G.A. Rouleau, G. Poptodorov, A. Jordanova, H. Rask-Andersen, L. Kluwe, V. Mautner, M. Sainio, G. Hung, T. Mathiesen, C. Moller, S.M. Pulst, H. Harder, A. Heiberg, M. Honda, M. Niimura, S. Sahlen, E. Blennow, D.G. Albertson, D. Pinkel, and J.P. Dumanski. 2001. High resolution deletion analysis of constitutional DNA from neurofibromatosis type 2 (NF2) patients using microarray-CGH.  Hum Mol Genet 10:271-282.

15. 	Hand, J.L., V.V. Michels, M.J. Marinello, R.P. Ketterling, and S.M. Jalal. 2000. Inherited interstitial deletion of chromosomes 5p and 16q without apparent phenotypic effect: further confirmation.  Prenat Diagn JID - 8106540 20:144-148.

16. 	Solinas-Toldo, S., S. Lampel, S. Stilgenbauer, J. Nickolenko, A. Benner, H. Dohner, T. Cremer, and P. Lichter. 1997. Matrix-based comparative genomic hybridization: biochips to screen for genomic imbalances.  Genes Chromosomes Cancer 20:399-407.

17. 	Shaw-Smith, C., R. Redon, L. Rickman, M. Rio, L. Willatt, H. Fiegler, H. Firth, D. Sanlaville, R. Winter, L. Colleaux, M. Bobrow, and N.P. Carter. 2004. Microarray based comparative genomic hybridisation (array-CGH) detects submicroscopic chromosomal deletions and duplications in patients with learning disability/mental retardation and dysmorphic features.  J Med Genet JID - 2985087R 41:241-248.

18. 	Albertson, D.G., B. Ylstra, R. Segraves, C. Collins, S.H. Dairkee, D. Kowbel, W.L. Kuo, J.W. Gray, and D. Pinkel. 2000. Quantitative mapping of amplicon structure by array CGH identifies CYP24 as a candidate oncogene.  Nat Genet JID - 9216904 25:144-146.

19. 	Lucito, R., J. Healy, J. Alexander, A. Reiner, D. Esposito, M. Chi, L. Rodgers, A. Brady, J. Sebat, J. Troge, J.A. West, S. Rostan, K.C. Nguyen, S. Powers, K.Q. Ye, A. Olshen, E. Venkatraman, L. Norton, and M. Wigler. 2003. Representational oligonucleotide microarray analysis: a high-resolution method to detect genome copy number variation.  Genome Res JID - 9518021 13:2291-2305.

20. 	Rosenberg, M.J., C. Killoran, L. Dziadzio, S. Chang, D.L. Stone, J. Meck, D. Aughton, L.M. Bird, J. Bodurtha, S.B. Cassidy, J.M.J. Graham, A. Grix, A.E. Guttmacher, L. Hudgins, C. Kozma, R.C. Michaelis, R. Pauli, K.F. Peters, K.N. Rosenbaum, C.J. Tifft, D. Wargowski, M.S. Williams, and L.G. Biesecker. 2001. Scanning for telomeric deletions and duplications and uniparental disomy using genetic markers in 120 children with malformations.  Hum Genet JID - 7613873 109:311-318.

21. 	Slavotinek, A., M. Rosenberg, S. Knight, L. Gaunt, W. Fergusson, C. Killoran, J. Clayton-Smith, H. Kingston, R.H. Campbell, J. Flint, D. Donnai, and L. Biesecker. 1999. Screening for submicroscopic chromosome rearrangements in children with idiopathic mental retardation using microsatellite markers for the chromosome telomeres.  J Med Genet 36:405-411.

22. 	Starke, H., A. Nietzel, A. Weise, A. Heller, K. Mrasek, B. Belitz, C. Kelbova, M. Volleth, B. Albrecht, B. Mitulla, R. Trappe, I. Bartels, S. Adolph, A. Dufke, S. Singer, M. Stumm, R.D. Wegner, J. Seidel, A. Schmidt, A. Kuechler, I. Schreyer, U. Claussen, F. von Eggeling, and T. Liehr. 2003. Small supernumerary marker chromosomes (SMCs): genotype-phenotype correlation and classification.  Hum Genet JID - 7613873 114:51-67.

23. 	Riegel, M., A. Baumer, M. Jamar, K. Delbecque, C. Herens, A. Verloes, and A. Schinzel. 2001. Submicroscopic terminal deletions and duplications in retarded patients with unclassified malformation syndromes.  Hum Genet 109:286-294.

24. 	Knight, S.J., R. Regan, A. Nicod, S.W. Horsley, L. Kearney, T. Homfray, R.M. Winter, P. Bolton, and J. Flint. 1999. Subtle chromosomal rearrangements in children with unexplained mental retardation.  Lancet 354:1676-1681.

25. 	Flint, J., A.O. Wilkie, V.J. Buckle, R.M. Winter, A.J. Holland, and H.E. McDermid. 1995. The detection of subtelomeric chromosomal rearrangements in idiopathic mental retardation.  Nat Genet 9:132-140.

26. 	Schoumans, J., B.M. Anderlid, E. Blennow, B.T. Teh, and M. Nordenskjold. 2004. The performance of CGH array for the detection of cryptic constitutional chromosome imbalances.  J Med Genet JID - 2985087R 41:198-202.

27.	Van Buggenhout G, Melotte C, Dutta D et al. The mild wolf-hirschhorn syndrome: microarray CGH analysis of atypical 4p16.3 deletions enables refinement of the genotype-phenotype map. J Med Genet 2004; in press.

Figures

## Figure 1:	(A-C) Chromosome 9 array CGH ratio profiles  using DNA from a patient with a duplication 9q22.33→31.1 using different ratios DNA/Cot-1 (indicated at each panel). For each panel, the X axis represents the clones ordered from the 9p telomere to the 9q telomere according to their clone position in the February 2004 Ensembl freeze.  The Y axis marks log2  transformed intensity ratios at each locus.

## Figure 2: Chromosome 9 array CGH ratio profiles  using DNA from a patient with a duplication 9q22.33→31.1  using 200 pmol DNA/100 µg Cot-1.  For each panel, the X axis represents the clones ordered from the 9p telomere to the 9q telomere according to their clone position in the February 2004 Ensembl freeze.  The Y axis marks the hybridization ratio plotted on a log2 scale.  (A) and (B) represent the analysis of the same sample hybridization but with inversed dye labeling.  Panel (C) represents the combined results of (A) and (B). In panel C the  red lines indicates the lower thresholds for clone deletion or duplication ( 4 x SD) and the green line the upper threshold (log2(3/2) - 2 x SD).  The Y axis shows the log2  transformed intensity ratios at each locus.

## Figure 3: Array CGH ratio profiles using DNA from a female trisomy 13 cell line (A and C) and a mixture containing 20% DNA from the trisomy 13 cell line and 80% DNA from a normal female cell line (B and D) versus DNA from a normal male cell line.  Panels A and B provide a genome wide view. Clones are ordered from the short arm telomere to  the long arm telomere and chromosomes are ordered from chromosomes1-9, X,Y and 10-22.  Panels C and D show a partial karyotype enlarging the ratio profiles obtained for chromosomes 12, 13 and 14.  Red lines indicate the lower thresholds for clone deletion or duplication ( 4 x SD) and the green line the upper threshold (log2(3/2) - 2 x SD).  The Y axis shows the log2  transformed intensity ratios at each locus.